## RTD Data
ENTER YOUR RTD VALUES HERE (BLUE FIELDS) | [9 empty] | SOFTWARE SENSOR SETTING RESULTS HERE | [7 empty] | CONNECTION DIAGRAM
Constants | [9 empty] | Recommend Using Sensitivity for 20°C
R0 @ 0° C | 100 | ohms | VEXC | 5 | V | iRTD 0°C | 0.49505 | mA | 100 | Sensitivity for T = 20°C |  |  |  | Sensitivity for T = Mid-Range
R0 Custom | 0 | ohms | TempLow | -50 | °C | TempRange | 550 | °C |  | Tempambient |  | 20 | °C | TempMid |  | 280 | °C
R0 Tolerance | 0 | ohms | TempHigh | 500 | °C |  |  |  |  | VRTD @ 20°C |  | 53.3203 | mV | VRTD @ Mid |  | 100.3929 | mV
Adjustable Variables | [9 empty] | Sensitivity20°C |  | 0.19013 | mV/°C | Sensitivity @Mid |  | 0.17212 | mV/°C
R2_0.5mA_calc | 4950 | ohms |  | R2_des_std_value | 5000 | ohms |  |  |  | Initial EU |  | -260.4421 | °C | Initial EU |  | -303.2585 | °C
R1_0.5mA_calc | 4950 | ohms |  | R1_des_std_value | 5000 | ohms |  |  |  | Desired Range/Capacity |  | 550 | °C | Desired Range/Capacity |  | 550 | °C
Standard Temperature Coefficient | [9 empty] | ∆T Information
Rcoeff | 0.00385 | Ω/Ω/°C | [7 empty] | TLOW_adj |  | -50 | °C | TLOW_adj |  | -50 | °C
[10 empty] | TMID_adj |  | 280 | °C | TLOW_adj |  | 280 | °C
[10 empty] | THIGH_adj |  | 500 | °C | THIGH_adj |  | 500 | °C
[11 empty] | Tactual | Tmeas | % Error |  | Tactual | Tmeas | % Error
[10 empty] | TLOW_adj | -50 | -50.8594 | 1.7188 | TLOW_adj | -50 | -56.1796 | 12.3591
[10 empty] | TMID_adj | 280 | 267.6494 | -4.4109 | TLOW_adj | 280 | 295.6468 | 5.5882
[10 empty] | THIGH_adj | 500 | 303.5839 | -5.13 | THIGH_adj | 500 | 335.3403 | 4.7938
Other Sensor Settings
[10 empty] | Bridge Mode |  | Full Bridge |  | Proportional |  | NO
[10 empty] | Check Offset |  | YES |  | Remove Offset |  | NO
[10 empty] | Zero Method |  | Absolute Zero

## RTD Resistance @ Temperature
T/°C | From RTD Table | T/°C | From RTD Table | T/°C | From RTD Table | T/°C | CUSTOM
| PT100 |  | PT500 |  | PT1000 |  | 0
-200 | 18.52 | -200 | 92.6 | -200 | 185.2 | -200 | 0
-190 | 22.83 | -190 | 114.15 | -190 | 228.3 | -190 | 0
-180 | 27.1 | -180 | 135.5 | -180 | 271 | -180 | 0
-170 | 31.34 | -170 | 156.7 | -170 | 313.4 | -170 | 0
-160 | 35.54 | -160 | 177.7 | -160 | 355.4 | -160 | 0
-150 | 39.72 | -150 | 198.6 | -150 | 397.2 | -150 | 0
-140 | 43.88 | -140 | 219.4 | -140 | 438.8 | -140 | 0
-130 | 48 | -130 | 240 | -130 | 480 | -130 | 0
-120 | 52.11 | -120 | 260.55 | -120 | 521.1 | -120 | 0
-110 | 56.19 | -110 | 280.95 | -110 | 561.9 | -110 | 0
-100 | 60.26 | -100 | 301.3 | -100 | 602.6 | -100 | 0
-90 | 64.3 | -90 | 321.5 | -90 | 643 | -90 | 0
-80 | 68.33 | -80 | 341.65 | -80 | 683.3 | -80 | 0
-70 | 72.33 | -70 | 361.65 | -70 | 723.3 | -70 | 0
-60 | 76.33 | -60 | 381.65 | -60 | 763.3 | -60 | 0
-50 | 80.31 | -50 | 401.55 | -50 | 803.1 | -50 | 0
-40 | 84.27 | -40 | 421.35 | -40 | 842.7 | -40 | 0
-30 | 88.22 | -30 | 441.1 | -30 | 882.2 | -30 | 0
-20 | 92.16 | -20 | 460.8 | -20 | 921.6 | -20 | 0
-10 | 96.09 | -10 | 480.45 | -10 | 960.9 | -10 | 0
0 | 100 | 0 | 500 | 0 | 1000 | 0 | 0
10 | 103.9 | 10 | 519.5 | 10 | 1039 | 10 | 0
20 | 107.79 | 20 | 538.95 | 20 | 1077.9 | 20 | 0
30 | 111.67 | 30 | 558.35 | 30 | 1116.7 | 30 | 0
40 | 115.54 | 40 | 577.7 | 40 | 1155.4 | 40 | 0
50 | 119.4 | 50 | 597 | 50 | 1194 | 50 | 0
60 | 123.24 | 60 | 616.2 | 60 | 1232.4 | 60 | 0
70 | 127.08 | 70 | 635.4 | 70 | 1270.8 | 70 | 0
80 | 130.9 | 80 | 654.5 | 80 | 1309 | 80 | 0
90 | 134.71 | 90 | 673.55 | 90 | 1347.1 | 90 | 0
100 | 138.51 | 100 | 692.55 | 100 | 1385.1 | 100 | 0
110 | 142.29 | 110 | 711.45 | 110 | 1422.9 | 110 | 0
120 | 146.07 | 120 | 730.35 | 120 | 1460.7 | 120 | 0
130 | 149.83 | 130 | 749.15 | 130 | 1498.3 | 130 | 0
140 | 153.58 | 140 | 767.9 | 140 | 1535.8 | 140 | 0
150 | 157.33 | 150 | 786.65 | 150 | 1573.3 | 150 | 0
160 | 161.05 | 160 | 805.25 | 160 | 1610.5 | 160 | 0
170 | 164.77 | 170 | 823.85 | 170 | 1647.7 | 170 | 0
180 | 168.48 | 180 | 842.4 | 180 | 1684.8 | 180 | 0
190 | 172.17 | 190 | 860.85 | 190 | 1721.7 | 190 | 0
200 | 175.86 | 200 | 879.3 | 200 | 1758.6 | 200 | 0
210 | 179.53 | 210 | 897.65 | 210 | 1795.3 | 210 | 0
220 | 183.19 | 220 | 915.95 | 220 | 1831.9 | 220 | 0
230 | 186.84 | 230 | 934.2 | 230 | 1868.4 | 230 | 0
240 | 190.47 | 240 | 952.35 | 240 | 1904.7 | 240 | 0
250 | 194.1 | 250 | 970.5 | 250 | 1941 | 250 | 0
260 | 197.71 | 260 | 988.55 | 260 | 1977.1 | 260 | 0
270 | 201.31 | 270 | 1006.55 | 270 | 2013.1 | 270 | 0
280 | 204.9 | 280 | 1024.5 | 280 | 2049 | 280 | 0
290 | 208.48 | 290 | 1042.4 | 290 | 2084.8 | 290 | 0
300 | 212.05 | 300 | 1060.25 | 300 | 2120.5 | 300 | 0
310 | 215.61 | 310 | 1078.05 | 310 | 2156.1 | 310 | 0
320 | 219.15 | 320 | 1095.75 | 320 | 2191.5 | 320 | 0

## PT100 Tables
Equations | [9 empty] | Measured Values
VRTD mV = RRTD * iRTD (mA) |  |  |  |  | MINIMUMS | VRDT_0.5mA = 0.5mA * R0 |  | 0.05 | V | VRTD @ 0°C | 49.505 | mV
| iRTD = VEXC / (R1 + R2 + RRTD) |  |  |  |  | Vdrop = VEXC - VRTD_0.5mA |  | 4.95 | V | VRTD @ 20°C | 53.3203 | mV
|  | VEXC | 5 | V |  | @ R1 = R2 |  |  |  | Sensitivity20°C | 0.19013 | mV/°C
|  | R1 | 5000 | ohms |  |  | VR = Vdrop / 2 | 2.475 | V | TempMid | 280 | °C
|  | R2 | 5000 | ohms |  |  | R2 = VR/0.5mA | 4950 | ohms | VRTD @ Mid | 100.3929 | mV
SBASE = VRTD 10°C / ∆T | [6 empty] | R1 = R2 | 4950 | ohms | Sensitivity @Mid | 0.17212 | mV/°C
ST = (SBASE @(T+10) - SBASE @(T-10)) / ∆T(T-10) to (T+10)
T/°C | RRTD | iRTD (A) | iRTD (mA) | VRTD (mV) | ∆VRTD 10°C (mV) | ∆V20°C (mV) | SBASE (mV/°C) | ST (mV/°C) |  | Tcalc for ST @ 20°C | ErrorT_calc @ 20°C (%) |  | ErrorT_calc @ Mid (%)
Tcalc @ Smid
-200 | 18.52 | 0.00049908 | 0.49908 | 9.2429 | 2.1461 | -44.0774 | 0.21461 | 0.21461 |  | -211.7615 | 5.8807 | -233.9128 | 16.9564
-190 | 22.83 | 0.00049886 | 0.49886 | 11.389 | 2.1461 | -41.9313 | 0.21461 | 0.21352 |  | -200.4738 | 5.5125 | -221.4444 | 16.5497
-180 | 27.1 | 0.00049865 | 0.49865 | 13.5134 | 2.1244 | -39.8069 | 0.21244 | 0.2116 |  | -189.3005 | 5.1669 | -209.1022 | 16.1679
-170 | 31.34 | 0.00049844 | 0.49844 | 15.621 | 2.1077 | -37.6992 | 0.21077 | 0.20968 |  | -178.2151 | 4.8324 | -196.8572 | 15.7984
-160 | 35.54 | 0.00049823 | 0.49823 | 17.7071 | 2.086 | -35.6132 | 0.2086 | 0.20802 |  | -167.2434 | 4.5271 | -184.7379 | 15.4612
-150 | 39.72 | 0.00049802 | 0.49802 | 19.7814 | 2.0744 | -33.5388 | 0.20744 | 0.20685 |  | -156.3332 | 4.2221 | -172.6864 | 15.1243
-140 | 43.88 | 0.00049782 | 0.49782 | 21.8441 | 2.0627 | -31.4761 | 0.20627 | 0.2052 |  | -145.4841 | 3.9172 | -160.7025 | 14.7875
-130 | 48 | 0.00049761 | 0.49761 | 23.8854 | 2.0412 | -29.4349 | 0.20412 | 0.20379 |  | -134.7483 | 3.6525 | -148.8436 | 14.4951
-120 | 52.11 | 0.00049741 | 0.49741 | 25.9199 | 2.0346 | -27.4003 | 0.20346 | 0.20263 |  | -124.0472 | 3.3727 | -137.0232 | 14.186
-110 | 56.19 | 0.00049721 | 0.49721 | 27.938 | 2.0181 | -25.3822 | 0.20181 | 0.20148 |  | -113.433 | 3.1209 | -125.2986 | 13.9078
-100 | 60.26 | 0.00049701 | 0.49701 | 29.9495 | 2.0115 | -23.3707 | 0.20115 | 0.20033 |  | -102.8533 | 2.8533 | -113.6122 | 13.6122
-90 | 64.3 | 0.00049681 | 0.49681 | 31.9446 | 1.9951 | -21.3757 | 0.19951 | 0.19918 |  | -92.3601 | 2.6223 | -102.0214 | 13.3571
-80 | 68.33 | 0.00049661 | 0.49661 | 33.9331 | 1.9885 | -19.3871 | 0.19885 | 0.19804 |  | -81.9012 | 2.3765 | -90.4684 | 13.0856
-70 | 72.33 | 0.00049641 | 0.49641 | 35.9053 | 1.9722 | -17.415 | 0.19722 | 0.19714 |  | -71.5284 | 2.1835 | -79.0107 | 12.8724
-60 | 76.33 | 0.00049621 | 0.49621 | 37.8759 | 1.9706 | -15.4444 | 0.19706 | 0.19649 |  | -61.1639 | 1.9399 | -67.562 | 12.6033
-50 | 80.31 | 0.00049602 | 0.49602 | 39.8351 | 1.9592 | -13.4852 | 0.19592 | 0.19535 |  | -50.8594 | 1.7188 | -56.1796 | 12.3591
-40 | 84.27 | 0.00049582 | 0.49582 | 41.7829 | 1.9478 | -11.5374 | 0.19478 | 0.19446 |  | -40.6148 | 1.5369 | -44.8633 | 12.1581
-30 | 88.22 | 0.00049563 | 0.49563 | 43.7243 | 1.9414 | -9.596 | 0.19414 | 0.19382 |  | -30.404 | 1.3466 | -33.5844 | 11.9479
-20 | 92.16 | 0.00049543 | 0.49543 | 45.6592 | 1.9349 | -7.6611 | 0.19349 | 0.19317 |  | -20.227 | 1.135 | -22.3428 | 11.7142
-10 | 96.09 | 0.00049524 | 0.49524 | 47.5877 | 1.9285 | -5.7325 | 0.19285 | 0.19229 |  | -10.0838 | 0.83773 | -11.1386 | 11.3859
0 | 100 | 0.00049505 | 0.49505 | 49.505 | 1.9172 | -3.8153 | 0.19172 | 0.1914 |  | 0 | 0 | 0 | 0 | 0°C
10 | 103.9 | 0.00049486 | 0.49486 | 51.4158 | 1.9108 | -1.9045 | 0.19108 | 0.19077 |  | 10.0502 | 0.50209 | 11.1015 | 11.0151
20 | 107.79 | 0.00049467 | 0.49467 | 53.3203 | 1.9045 | 0 | 0.19045 | 0.19013 |  | 20.0669 | 0.33461 | 22.166 | 10.8301 | AMBIENT
30 | 111.67 | 0.00049448 | 0.49448 | 55.2184 | 1.8981 | 1.8981 | 0.18981 | 0.18949 |  | 30.0502 | 0.16736 | 33.1936 | 10.6454
40 | 115.54 | 0.00049429 | 0.49429 | 57.1101 | 1.8918 | 3.7899 | 0.18918 | 0.18886 |  | 40.0001 | 0.00034112 | 44.1843 | 10.4609
50 | 119.4 | 0.0004941 | 0.4941 | 58.9956 | 1.8854 | 5.6753 | 0.18854 | 0.18798 |  | 49.9168 | -0.16646 | 55.1383 | 10.2766
60 | 123.24 | 0.00049391 | 0.49391 | 60.8698 | 1.8742 | 7.5496 | 0.18742 | 0.18735 |  | 59.7745 | -0.37579 | 66.0272 | 10.0454
70 | 127.08 | 0.00049373 | 0.49373 | 62.7427 | 1.8728 | 9.4224 | 0.18728 | 0.18672 |  | 69.6248 | -0.536 | 76.9079 | 9.8684
80 | 130.9 | 0.00049354 | 0.49354 | 64.6043 | 1.8617 | 11.2841 | 0.18617 | 0.18585 |  | 79.4164 | -0.72955 | 87.7237 | 9.6546
90 | 134.71 | 0.00049335 | 0.49335 | 66.4597 | 1.8554 | 13.1395 | 0.18554 | 0.18523 |  | 89.1749 | -0.91674 | 98.5031 | 9.4479
100 | 138.51 | 0.00049317 | 0.49317 | 68.3089 | 1.8491 | 14.9886 | 0.18491 | 0.18436 |  | 98.9006 | -1.0994 | 109.2461 | 9.2461
110 | 142.29 | 0.00049299 | 0.49299 | 70.1469 | 1.838 | 16.8266 | 0.1838 | 0.18373 |  | 108.5678 | -1.302 | 119.9246 | 9.0223
120 | 146.07 | 0.0004928 | 0.4928 | 71.9835 | 1.8367 | 18.6633 | 0.18367 | 0.18311 |  | 118.2279 | -1.4768 | 130.5951 | 8.8292
130 | 149.83 | 0.00049262 | 0.49262 | 73.8091 | 1.8256 | 20.4889 | 0.18256 | 0.18225 |  | 127.8297 | -1.6695 | 141.2013 | 8.6164
140 | 153.58 | 0.00049244 | 0.49244 | 75.6285 | 1.8194 | 22.3082 | 0.18194 | 0.18187 |  | 137.3988 | -1.858 | 151.7714 | 8.4082
150 | 157.33 | 0.00049226 | 0.49226 | 77.4465 | 1.818 | 24.1263 | 0.1818 | 0.18101 |  | 146.9609 | -2.026 | 162.3338 | 8.2225
160 | 161.05 | 0.00049208 | 0.49208 | 79.2487 | 1.8022 | 25.9284 | 0.18022 | 0.18015 |  | 156.4396 | -2.2253 | 172.8039 | 8.0024
170 | 164.77 | 0.0004919 | 0.4919 | 81.0495 | 1.8008 | 27.7293 | 0.18008 | 0.17978 |  | 165.9113 | -2.4051 | 183.2664 | 7.8038
180 | 168.48 | 0.00049172 | 0.49172 | 82.8442 | 1.7947 | 29.524 | 0.17947 | 0.17892 |  | 175.3506 | -2.583 | 193.6931 | 7.6073
190 | 172.17 | 0.00049154 | 0.49154 | 84.628 | 1.7837 | 31.3077 | 0.17837 | 0.17831 |  | 184.7322 | -2.7725 | 204.0561 | 7.3979
200 | 175.86 | 0.00049136 | 0.49136 | 86.4104 | 1.7824 | 33.0901 | 0.17824 | 0.1777 |  | 194.107 | -2.9465 | 214.4116 | 7.2058
210 | 179.53 | 0.00049118 | 0.49118 | 88.1819 | 1.7715 | 34.8616 | 0.17715 | 0.17684 |  | 203.4243 | -3.1313 | 224.7035 | 7.0016
220 | 183.19 | 0.00049101 | 0.49101 | 89.9473 | 1.7654 | 36.627 | 0.17654 | 0.17623 |  | 212.7095 | -3.3139 | 234.9599 | 6.8
230 | 186.84 | 0.00049083 | 0.49083 | 91.7066 | 1.7593 | 38.3863 | 0.17593 | 0.17539 |  | 221.9626 | -3.4945 | 245.181 | 6.6004
240 | 190.47 | 0.00049065 | 0.49065 | 93.455 | 1.7484 | 40.1347 | 0.17484 | 0.17478 |  | 231.1585 | -3.6839 | 255.3388 | 6.3912
250 | 194.1 | 0.00049048 | 0.49048 | 95.2021 | 1.7472 | 41.8819 | 0.17472 | 0.17417 |  | 240.3479 | -3.8609 | 265.4894 | 6.1958
260 | 197.71 | 0.00049031 | 0.49031 | 96.9384 | 1.7363 | 43.6182 | 0.17363 | 0.17333 |  | 249.4801 | -4.0461 | 275.5769 | 5.9911
270 | 201.31 | 0.00049013 | 0.49013 | 98.6687 | 1.7303 | 45.3484 | 0.17303 | 0.17273 |  | 258.5806 | -4.2294 | 285.6294 | 5.7887
280 | 204.9 | 0.00048996 | 0.48996 | 100.3929 | 1.7242 | 47.0727 | 0.17242 | 0.17212 |  | 267.6494 | -4.4109 | 295.6468 | 5.5882
290 | 208.48 | 0.00048979 | 0.48979 | 102.1112 | 1.7182 | 48.7909 | 0.17182 | 0.17152 |  | 276.6866 | -4.5908 | 305.6294 | 5.3894
300 | 212.05 | 0.00048962 | 0.48962 | 103.8234 | 1.7122 | 50.5032 | 0.17122 | 0.17092 |  | 285.6922 | -4.7693 | 315.5771 | 5.1924
310 | 215.61 | 0.00048945 | 0.48945 | 105.5297 | 1.7063 | 52.2094 | 0.17063 | 0.17009 |  | 294.6664 | -4.9463 | 325.49 | 4.9968
320 | 219.15 | 0.00048928 | 0.48928 | 107.2252 | 1.6955 | 53.9049 | 0.16955 | 0.16955 |  | 303.5839 | -5.13 | 335.3403 | 4.7938

## PT500 Tables
Equations | [9 empty] | Measured Values
VRTD mV = RRTD * iRTD (mA) |  |  |  |  | MINIMUMS | VRDT_0.5mA = 0.5mA * R0 |  | 0.25 | V | VRTD @ 0°C | 238.0952 | mV
| iRTD = VEXC / (R1 + R2 + RRTD) |  |  |  |  | Vdrop = VEXC - VRTD_0.5mA |  | 4.75 | V | VRTD @ 20°C | 255.6944 | mV
|  | VEXC | 5 | V |  | @ R1 = R2 |  |  |  | Sensitivity20°C | 0.87446 | mV/°C
|  | R1 | 5000 | ohms |  |  | VR = Vdrop / 2 | 2.375 | V | TempMid | 280 | °C
|  | R2 | 5000 | ohms |  |  | R2 = VR/0.5mA | 4750 | ohms | VRTD @ Mid | 464.6469 | mV
SBASE = VRTD 10°C / ∆T | [6 empty] | R1 = R2 | 4750 | ohms | Sensitivity @Mid | 0.73742 | mV/°C
ST = (SBASE @(T+10) - SBASE @(T-10)) / ∆T(T-10) to (T+10)
T/°C | RRTD | iRTD (A) | iRTD (mA) | VRTD (mV) | ∆VRTD 10°C (mV) | ∆V20°C (mV) | SBASE (mV/°C) | ST (mV/°C) |  | Tcalc for ST @ 20°C | ErrorT_calc @ 20°C (%) |  | ErrorT_calc @ Mid (%)
Tcalc @ Smid
-200 | 92.6 | 0.00049541 | 0.49541 | 45.8752 | 10.5556 | -209.8192 | 1.0556 | 1.0556 |  | -219.8157 | 9.9079 | -260.6659 | 30.3329
-190 | 114.15 | 0.00049436 | 0.49436 | 56.4308 | 10.5556 | -199.2635 | 1.0556 | 1.0485 |  | -207.7447 | 9.3393 | -246.3516 | 29.6587
-180 | 135.5 | 0.00049332 | 0.49332 | 66.8443 | 10.4134 | -188.8501 | 1.0413 | 1.0355 |  | -195.8363 | 8.7979 | -232.2301 | 29.0167
-170 | 156.7 | 0.00049229 | 0.49229 | 77.1412 | 10.2969 | -178.5532 | 1.0297 | 1.0227 |  | -184.0611 | 8.2712 | -218.2667 | 28.3922
-160 | 177.7 | 0.00049127 | 0.49127 | 87.2987 | 10.1575 | -168.3957 | 1.0158 | 1.0113 |  | -172.4453 | 7.7783 | -204.4923 | 27.8077
-150 | 198.6 | 0.00049026 | 0.49026 | 97.3663 | 10.0676 | -158.328 | 1.0068 | 1.0023 |  | -160.9324 | 7.2883 | -190.8398 | 27.2265
-140 | 219.4 | 0.00048927 | 0.48927 | 107.3449 | 9.9785 | -148.3495 | 0.99785 | 0.99106 |  | -149.5213 | 6.8009 | -177.3081 | 26.6486
-130 | 240 | 0.00048828 | 0.48828 | 117.1875 | 9.8426 | -138.5069 | 0.98426 | 0.9811 |  | -138.2656 | 6.3582 | -163.9606 | 26.1236
-120 | 260.55 | 0.0004873 | 0.4873 | 126.9669 | 9.7794 | -128.7275 | 0.97794 | 0.97244 |  | -127.0823 | 5.9019 | -150.699 | 25.5825
-110 | 280.95 | 0.00048634 | 0.48634 | 136.6362 | 9.6693 | -119.0581 | 0.96693 | 0.96384 |  | -116.0248 | 5.4771 | -137.5866 | 25.0788
-100 | 301.3 | 0.00048538 | 0.48538 | 146.2437 | 9.6075 | -109.4507 | 0.96075 | 0.95533 |  | -105.0381 | 5.0381 | -124.5581 | 24.5581
-90 | 321.5 | 0.00048443 | 0.48443 | 155.7429 | 9.4992 | -99.9515 | 0.94992 | 0.94689 |  | -94.1751 | 4.639 | -111.6765 | 24.085
-80 | 341.65 | 0.00048348 | 0.48348 | 165.1816 | 9.4387 | -90.5128 | 0.94387 | 0.93854 |  | -83.3814 | 4.2267 | -98.8768 | 23.596
-70 | 361.65 | 0.00048255 | 0.48255 | 174.5137 | 9.3321 | -81.1806 | 0.93321 | 0.93142 |  | -72.7095 | 3.8707 | -86.2217 | 23.1738
-60 | 381.65 | 0.00048162 | 0.48162 | 183.8099 | 9.2962 | -71.8845 | 0.92962 | 0.92552 |  | -62.0787 | 3.4645 | -73.6153 | 22.6922
-50 | 401.55 | 0.0004807 | 0.4807 | 193.0241 | 9.2142 | -62.6702 | 0.92142 | 0.91736 |  | -51.5417 | 3.0833 | -61.1201 | 22.2402
-40 | 421.35 | 0.00047978 | 0.47978 | 202.1571 | 9.133 | -53.5372 | 0.9133 | 0.91042 |  | -41.0975 | 2.7438 | -48.735 | 21.8375
-30 | 441.1 | 0.00047888 | 0.47888 | 211.2325 | 9.0754 | -44.4618 | 0.90754 | 0.90469 |  | -30.7192 | 2.3973 | -36.428 | 21.4266
-20 | 460.8 | 0.00047797 | 0.47797 | 220.2508 | 9.0183 | -35.4435 | 0.90183 | 0.899 |  | -20.4062 | 2.031 | -24.1984 | 20.9922
-10 | 480.45 | 0.00047708 | 0.47708 | 229.2125 | 8.9616 | -26.4819 | 0.89616 | 0.89222 |  | -10.158 | 1.5799 | -12.0457 | 20.4573
0 | 500 | 0.00047619 | 0.47619 | 238.0952 | 8.8828 | -17.5991 | 0.88828 | 0.88549 |  | 0 | 0 | 0 | 0 | 0°C
10 | 519.5 | 0.00047531 | 0.47531 | 246.9224 | 8.8271 | -8.772 | 0.88271 | 0.87996 |  | 10.0944 | 0.94396 | 11.9703 | 19.7032
20 | 538.95 | 0.00047443 | 0.47443 | 255.6944 | 8.772 | 0 | 0.8772 | 0.87446 |  | 20.1257 | 0.62849 | 23.8658 | 19.3291 | AMBIENT
30 | 558.35 | 0.00047356 | 0.47356 | 264.4116 | 8.7172 | 8.7172 | 0.87172 | 0.86901 |  | 30.0944 | 0.31465 | 35.6871 | 18.9569
40 | 577.7 | 0.00047269 | 0.47269 | 273.0745 | 8.6629 | 17.3801 | 0.86629 | 0.8636 |  | 40.001 | 0.0024425 | 47.4347 | 18.5867
50 | 597 | 0.00047183 | 0.47183 | 281.6835 | 8.609 | 25.9891 | 0.8609 | 0.85712 |  | 49.8459 | -0.30816 | 59.1092 | 18.2184
60 | 616.2 | 0.00047098 | 0.47098 | 290.2168 | 8.5333 | 34.5225 | 0.85333 | 0.85179 |  | 59.6043 | -0.65944 | 70.6811 | 17.8018
70 | 635.4 | 0.00047013 | 0.47013 | 298.7194 | 8.5025 | 43.025 | 0.85025 | 0.84652 |  | 69.3275 | -0.96069 | 82.2112 | 17.4446
80 | 654.5 | 0.00046929 | 0.46929 | 307.1472 | 8.4278 | 51.4529 | 0.84278 | 0.84018 |  | 78.9653 | -1.2934 | 93.64 | 17.0501
90 | 673.55 | 0.00046845 | 0.46845 | 315.523 | 8.3757 | 59.8286 | 0.83757 | 0.83499 |  | 88.5435 | -1.6184 | 104.9982 | 16.6647
100 | 692.55 | 0.00046762 | 0.46762 | 323.847 | 8.324 | 68.1526 | 0.8324 | 0.82875 |  | 98.0625 | -1.9375 | 116.2863 | 16.2863
110 | 711.45 | 0.00046679 | 0.46679 | 332.0979 | 8.2509 | 76.4035 | 0.82509 | 0.82364 |  | 107.498 | -2.2746 | 127.4752 | 15.8865
120 | 730.35 | 0.00046597 | 0.46597 | 340.3197 | 8.2219 | 84.6254 | 0.82219 | 0.81858 |  | 116.9002 | -2.5832 | 138.6247 | 15.5206
130 | 749.15 | 0.00046515 | 0.46515 | 348.4694 | 8.1497 | 92.7751 | 0.81497 | 0.81247 |  | 126.2198 | -2.9078 | 149.6763 | 15.1356
140 | 767.9 | 0.00046434 | 0.46434 | 356.5691 | 8.0996 | 100.8747 | 0.80996 | 0.80856 |  | 135.4823 | -3.2269 | 160.6601 | 14.7572
150 | 786.65 | 0.00046354 | 0.46354 | 364.6406 | 8.0715 | 108.9462 | 0.80715 | 0.80254 |  | 144.7125 | -3.525 | 171.6056 | 14.4038
160 | 805.25 | 0.00046274 | 0.46274 | 372.6198 | 7.9792 | 116.9254 | 0.79792 | 0.79655 |  | 153.8373 | -3.8517 | 182.4261 | 14.0163
170 | 823.85 | 0.00046194 | 0.46194 | 380.5716 | 7.9518 | 124.8773 | 0.79518 | 0.79275 |  | 162.9307 | -4.1584 | 193.2094 | 13.6526
180 | 842.4 | 0.00046115 | 0.46115 | 388.4749 | 7.9033 | 132.7805 | 0.79033 | 0.78686 |  | 171.9686 | -4.4619 | 203.9269 | 13.2927
190 | 860.85 | 0.00046037 | 0.46037 | 396.3088 | 7.8339 | 140.6144 | 0.78339 | 0.78206 |  | 180.9271 | -4.7752 | 214.5503 | 12.9212
200 | 879.3 | 0.00045959 | 0.45959 | 404.1161 | 7.8073 | 148.4217 | 0.78073 | 0.7773 |  | 189.8553 | -5.0724 | 225.1376 | 12.5688
210 | 897.65 | 0.00045881 | 0.45881 | 411.8548 | 7.7388 | 156.1605 | 0.77388 | 0.77153 |  | 198.7051 | -5.3785 | 235.632 | 12.2057
220 | 915.95 | 0.00045805 | 0.45805 | 419.5466 | 7.6918 | 163.8523 | 0.76918 | 0.76684 |  | 207.5011 | -5.6813 | 246.0627 | 11.8467
230 | 934.2 | 0.00045728 | 0.45728 | 427.1917 | 7.6451 | 171.4974 | 0.76451 | 0.76115 |  | 216.2438 | -5.981 | 256.4301 | 11.4914
240 | 952.35 | 0.00045652 | 0.45652 | 434.7697 | 7.578 | 179.0754 | 0.7578 | 0.75654 |  | 224.9097 | -6.2876 | 266.7064 | 11.1277
250 | 970.5 | 0.00045577 | 0.45577 | 442.3226 | 7.5529 | 186.6282 | 0.75529 | 0.75197 |  | 233.5469 | -6.5813 | 276.9488 | 10.7795
260 | 988.55 | 0.00045502 | 0.45502 | 449.8091 | 7.4865 | 194.1148 | 0.74865 | 0.74639 |  | 242.1082 | -6.8815 | 287.1011 | 10.4235
270 | 1006.55 | 0.00045427 | 0.45427 | 457.2505 | 7.4413 | 201.5561 | 0.74413 | 0.74189 |  | 250.6178 | -7.1786 | 297.1921 | 10.0712
280 | 1024.5 | 0.00045354 | 0.45354 | 464.6469 | 7.3965 | 208.9526 | 0.73965 | 0.73742 |  | 259.0761 | -7.4728 | 307.2224 | 9.7223
290 | 1042.4 | 0.0004528 | 0.4528 | 471.9988 | 7.3519 | 216.3045 | 0.73519 | 0.73298 |  | 267.4835 | -7.7643 | 317.1921 | 9.3766
300 | 1060.25 | 0.00045207 | 0.45207 | 479.3065 | 7.3077 | 223.6122 | 0.73077 | 0.72857 |  | 275.8403 | -8.0532 | 327.102 | 9.034
310 | 1078.05 | 0.00045134 | 0.45134 | 486.5703 | 7.2638 | 230.8759 | 0.72638 | 0.72318 |  | 284.1469 | -8.3397 | 336.9522 | 8.6943
320 | 1095.75 | 0.00045062 | 0.45062 | 493.7701 | 7.1998 | 238.0758 | 0.71998 | 0.71998 |  | 292.3804 | -8.6311 | 346.7158 | 8.3487

## PT1000 Tables
Equations | [9 empty] | Measured Values
VRTD mV = RRTD * iRTD (mA) |  |  |  |  | MINIMUMS | VRDT_0.5mA = 0.5mA * R0 |  | 0.5 | V | VRTD @ 0°C | 454.5455 | mV
| iRTD = VEXC / (R1 + R2 + RRTD) |  |  |  |  | Vdrop = VEXC - VRTD_0.5mA |  | 4.5 | V | VRTD @ 20°C | 486.5092 | mV
|  | VEXC | 5 | V |  | @ R1 = R2 |  |  |  | Sensitivity20°C | 1.5829 | mV/°C
|  | R1 | 5000 | ohms |  |  | VR = Vdrop / 2 | 2.25 | V | TempMid | 280 | °C
|  | R2 | 5000 | ohms |  |  | R2 = VR/0.5mA | 4500 | ohms | VRTD @ Mid | 850.278 | mV
SBASE = VRTD 10°C / ∆T | [6 empty] | R1 = R2 | 4500 | ohms | Sensitivity @Mid | 1.2456 | mV/°C
ST = (SBASE @(T+10) - SBASE @(T-10)) / ∆T(T-10) to (T+10)
T/°C | RRTD | iRTD (A) | iRTD (mA) | VRTD (mV) | ∆VRTD 10°C (mV) | ∆V20°C (mV) | SBASE (mV/°C) | ST (mV/°C) |  | Tcalc for ST @ 20°C | ErrorT_calc @ 20°C (%) |  | ErrorT_calc @ Mid (%)
Tcalc @ Smid
-200 | 185.2 | 0.00049091 | 0.49091 | 90.9162 | 20.6859 | -395.5929 | 2.0686 | 2.0686 |  | -229.7224 | 14.8612 | -291.9393 | 45.9696
-190 | 228.3 | 0.00048884 | 0.48884 | 111.6021 | 20.6859 | -374.9071 | 2.0686 | 2.0504 |  | -216.6541 | 14.0285 | -275.3317 | 44.9114
-180 | 271 | 0.00048681 | 0.48681 | 131.9248 | 20.3227 | -354.5843 | 2.0323 | 2.0168 |  | -203.8153 | 13.2307 | -259.0156 | 43.8975
-170 | 313.4 | 0.00048481 | 0.48481 | 151.9383 | 20.0134 | -334.5709 | 2.0013 | 1.9838 |  | -191.1718 | 12.454 | -242.9478 | 42.9105
-160 | 355.4 | 0.00048284 | 0.48284 | 171.6013 | 19.663 | -314.9079 | 1.9663 | 1.9537 |  | -178.7497 | 11.7186 | -227.1614 | 41.9759
-150 | 397.2 | 0.0004809 | 0.4809 | 191.013 | 19.4117 | -295.4962 | 1.9412 | 1.9288 |  | -166.4864 | 10.9909 | -211.5768 | 41.0512
-140 | 438.8 | 0.00047898 | 0.47898 | 210.1774 | 19.1644 | -276.3318 | 1.9164 | 1.8997 |  | -154.3793 | 10.2709 | -196.1906 | 40.1361
-130 | 480 | 0.0004771 | 0.4771 | 229.0076 | 18.8302 | -257.5015 | 1.883 | 1.8734 |  | -142.4833 | 9.6025 | -181.0728 | 39.2868
-120 | 521.1 | 0.00047524 | 0.47524 | 247.6452 | 18.6376 | -238.864 | 1.8638 | 1.8498 |  | -130.709 | 8.9242 | -166.1096 | 38.4247
-110 | 561.9 | 0.0004734 | 0.4734 | 266.0033 | 18.3581 | -220.5059 | 1.8358 | 1.8265 |  | -119.1113 | 8.283 | -151.3709 | 37.6099
-100 | 602.6 | 0.00047158 | 0.47158 | 284.1756 | 18.1723 | -202.3336 | 1.8172 | 1.8037 |  | -107.631 | 7.631 | -136.7813 | 36.7813
-90 | 643 | 0.00046979 | 0.46979 | 302.0765 | 17.9009 | -184.4327 | 1.7901 | 1.7811 |  | -96.3221 | 7.0246 | -122.4095 | 36.0106
-80 | 683.3 | 0.00046802 | 0.46802 | 319.7982 | 17.7217 | -166.711 | 1.7722 | 1.759 |  | -85.1265 | 6.4081 | -108.1817 | 35.2271
-70 | 723.3 | 0.00046627 | 0.46627 | 337.2563 | 17.4581 | -149.2529 | 1.7458 | 1.7393 |  | -74.0973 | 5.8533 | -94.1655 | 34.5221
-60 | 763.3 | 0.00046454 | 0.46454 | 354.5846 | 17.3283 | -131.9246 | 1.7328 | 1.7221 |  | -63.1502 | 5.2503 | -80.2535 | 33.7558
-50 | 803.1 | 0.00046283 | 0.46283 | 371.6989 | 17.1143 | -114.8103 | 1.7114 | 1.7009 |  | -52.3383 | 4.6765 | -66.5133 | 33.0266
-40 | 842.7 | 0.00046114 | 0.46114 | 388.6025 | 16.9036 | -97.9067 | 1.6904 | 1.6821 |  | -41.6594 | 4.1485 | -52.9422 | 32.3556
-30 | 882.2 | 0.00045947 | 0.45947 | 405.3408 | 16.7384 | -81.1683 | 1.6738 | 1.6657 |  | -31.085 | 3.6166 | -39.5039 | 31.6796
-20 | 921.6 | 0.00045781 | 0.45781 | 421.9162 | 16.5754 | -64.593 | 1.6575 | 1.6495 |  | -20.6135 | 3.0675 | -26.1964 | 30.9818
-10 | 960.9 | 0.00045617 | 0.45617 | 438.3308 | 16.4146 | -48.1784 | 1.6415 | 1.6315 |  | -10.2436 | 2.436 | -13.0179 | 30.1792
0 | 1000 | 0.00045455 | 0.45455 | 454.5455 | 16.2147 | -31.9637 | 1.6215 | 1.6137 |  | 0 | 0 | 0 | 0 | 0°C
10 | 1039 | 0.00045294 | 0.45294 | 470.6042 | 16.0588 | -15.905 | 1.6059 | 1.5982 |  | 10.1451 | 1.4511 | 12.8928 | 28.9276
20 | 1077.9 | 0.00045135 | 0.45135 | 486.5092 | 15.905 | 0 | 1.5905 | 1.5829 |  | 20.1931 | 0.96525 | 25.662 | 28.3102 | AMBIENT
30 | 1116.7 | 0.00044977 | 0.44977 | 502.2624 | 15.7532 | 15.7532 | 1.5753 | 1.5678 |  | 30.1451 | 0.4837 | 38.3095 | 27.6982
40 | 1155.4 | 0.00044821 | 0.44821 | 517.8658 | 15.6034 | 31.3566 | 1.5603 | 1.553 |  | 40.0026 | 0.0063983 | 50.8367 | 27.0916
50 | 1194 | 0.00044667 | 0.44667 | 533.3214 | 15.4556 | 46.8122 | 1.5456 | 1.5363 |  | 49.7666 | -0.46672 | 63.2452 | 26.4904
60 | 1232.4 | 0.00044514 | 0.44514 | 548.5916 | 15.2702 | 62.0824 | 1.527 | 1.5218 |  | 59.4135 | -0.97742 | 75.5048 | 25.8414
70 | 1270.8 | 0.00044362 | 0.44362 | 563.7577 | 15.1661 | 77.2485 | 1.5166 | 1.5076 |  | 68.9947 | -1.4361 | 87.6809 | 25.2584
80 | 1309 | 0.00044213 | 0.44213 | 578.7426 | 14.9849 | 92.2334 | 1.4985 | 1.4915 |  | 78.4614 | -1.9232 | 99.7115 | 24.6394
90 | 1347.1 | 0.00044064 | 0.44064 | 593.5878 | 14.8452 | 107.0786 | 1.4845 | 1.4776 |  | 87.8399 | -2.4002 | 111.63 | 24.0333
100 | 1385.1 | 0.00043917 | 0.43917 | 608.2951 | 14.7073 | 121.7859 | 1.4707 | 1.462 |  | 97.1312 | -2.8688 | 123.4377 | 23.4377
110 | 1422.9 | 0.00043772 | 0.43772 | 622.8278 | 14.5328 | 136.3187 | 1.4533 | 1.4485 |  | 106.3122 | -3.3525 | 135.1053 | 22.823
120 | 1460.7 | 0.00043627 | 0.43627 | 637.2647 | 14.4369 | 150.7556 | 1.4437 | 1.4352 |  | 115.4327 | -3.8061 | 146.696 | 22.2466
130 | 1498.3 | 0.00043485 | 0.43485 | 651.5311 | 14.2664 | 165.0219 | 1.4266 | 1.4201 |  | 124.4455 | -4.2727 | 158.1497 | 21.6536
140 | 1535.8 | 0.00043343 | 0.43343 | 665.6669 | 14.1358 | 179.1577 | 1.4136 | 1.409 |  | 133.3758 | -4.7316 | 169.4986 | 21.0704
150 | 1573.3 | 0.00043203 | 0.43203 | 679.7111 | 14.0442 | 193.2019 | 1.4044 | 1.3943 |  | 142.2482 | -5.1679 | 180.7739 | 20.516
160 | 1610.5 | 0.00043064 | 0.43064 | 693.5532 | 13.8422 | 207.0441 | 1.3842 | 1.3798 |  | 150.9929 | -5.6294 | 191.8871 | 19.9295
170 | 1647.7 | 0.00042927 | 0.42927 | 707.307 | 13.7538 | 220.7978 | 1.3754 | 1.3692 |  | 159.6819 | -6.0695 | 202.9293 | 19.3702
180 | 1684.8 | 0.00042791 | 0.42791 | 720.9366 | 13.6296 | 234.4274 | 1.363 | 1.355 |  | 168.2924 | -6.5042 | 213.8718 | 18.8177
190 | 1721.7 | 0.00042656 | 0.42656 | 734.4071 | 13.4705 | 247.8979 | 1.3471 | 1.3428 |  | 176.8023 | -6.9461 | 224.6866 | 18.2561
200 | 1758.6 | 0.00042522 | 0.42522 | 747.7931 | 13.386 | 261.2839 | 1.3386 | 1.3308 |  | 185.2589 | -7.3705 | 235.4335 | 17.7168
210 | 1795.3 | 0.0004239 | 0.4239 | 761.0235 | 13.2304 | 274.5143 | 1.323 | 1.3171 |  | 193.6172 | -7.8013 | 246.0555 | 17.1693
220 | 1831.9 | 0.00042259 | 0.42259 | 774.136 | 13.1126 | 287.6268 | 1.3113 | 1.3054 |  | 201.901 | -8.2268 | 256.5829 | 16.6286
230 | 1868.4 | 0.00042129 | 0.42129 | 787.1322 | 12.9962 | 300.623 | 1.2996 | 1.2921 |  | 210.1114 | -8.6472 | 267.0169 | 16.0943
240 | 1904.7 | 0.00042 | 0.42 | 799.9782 | 12.8459 | 313.469 | 1.2846 | 1.2807 |  | 218.2268 | -9.0722 | 277.3302 | 15.5543
250 | 1941 | 0.00041873 | 0.41873 | 812.746 | 12.7678 | 326.2368 | 1.2768 | 1.2694 |  | 226.2929 | -9.4829 | 287.5809 | 15.0324
260 | 1977.1 | 0.00041746 | 0.41746 | 825.3667 | 12.6207 | 338.8576 | 1.2621 | 1.2565 |  | 234.266 | -9.8977 | 297.7134 | 14.5052
270 | 2013.1 | 0.00041621 | 0.41621 | 837.877 | 12.5102 | 351.3678 | 1.251 | 1.2456 |  | 242.1693 | -10.3077 | 307.7573 | 13.9842
280 | 2049 | 0.00041497 | 0.41497 | 850.278 | 12.401 | 363.7689 | 1.2401 | 1.2347 |  | 250.0037 | -10.713 | 317.7134 | 13.4691
290 | 2084.8 | 0.00041374 | 0.41374 | 862.5712 | 12.2931 | 376.062 | 1.2293 | 1.224 |  | 257.7699 | -11.1138 | 327.583 | 12.9596
300 | 2120.5 | 0.00041252 | 0.41252 | 874.7576 | 12.1865 | 388.2485 | 1.2186 | 1.2134 |  | 265.4686 | -11.5105 | 337.3669 | 12.4556
310 | 2156.1 | 0.00041132 | 0.41132 | 886.8387 | 12.0811 | 400.3295 | 1.2081 | 1.2012 |  | 273.1008 | -11.903 | 347.0661 | 11.9568
320 | 2191.5 | 0.00041012 | 0.41012 | 898.7819 | 11.9432 | 412.2728 | 1.1943 | 1.1943 |  | 280.646 | -12.2981 | 356.6547 | 11.4546

## PT-CUSTOM Tables
Equations | [9 empty] | Measured Values
VRTD mV = RRTD * iRTD (mA) |  |  |  |  | MINIMUMS | VRDT_0.5mA = 0.5mA * R0 |  | 0.5 | V | VRTD @ 0°C | 0 | mV
| iRTD = VEXC / (R1 + R2 + RRTD) |  |  |  |  | Vdrop = VEXC - VRTD_0.5mA |  | 4.5 | V | VRTD @ 20°C | 0 | mV
|  | VEXC | 5 | V |  | @ R1 = R2 |  |  |  | Sensitivity20°C | 0 | mV/°C
|  | R1 | 5000 | ohms |  |  | VR = Vdrop / 2 | 2.25 | V | TempMid | 280 | °C
|  | R2 | 5000 | ohms |  |  | R2 = VR/0.5mA | 4500 | ohms | VRTD @ Mid | 0 | mV
SBASE = VRTD 10°C / ∆T | [6 empty] | R1 = R2 | 4500 | ohms | Sensitivity @Mid | 0 | mV/°C
ST = (SBASE @(T+10) - SBASE @(T-10)) / ∆T(T-10) to (T+10)
T/°C | RRTD | iRTD (A) | iRTD (mA) | VRTD (mV) | ∆VRTD 10°C (mV) | ∆V20°C (mV) | SBASE (mV/°C) | ST (mV/°C) |  | Tcalc for ST @ 20°C | ErrorT_calc @ 20°C (%) |  | ErrorT_calc @ Mid (%)
Tcalc @ Smid
-200 | 0 | 0.0005 | 0.5 | 0 | 0 | 0 | 0 | 0 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
-190 | 0 | 0.0005 | 0.5 | 0 | 0 | 0 | 0 | 0 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
-180 | 0 | 0.0005 | 0.5 | 0 | 0 | 0 | 0 | 0 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
-170 | 0 | 0.0005 | 0.5 | 0 | 0 | 0 | 0 | 0 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
-160 | 0 | 0.0005 | 0.5 | 0 | 0 | 0 | 0 | 0 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
-150 | 0 | 0.0005 | 0.5 | 0 | 0 | 0 | 0 | 0 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
-140 | 0 | 0.0005 | 0.5 | 0 | 0 | 0 | 0 | 0 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
-130 | 0 | 0.0005 | 0.5 | 0 | 0 | 0 | 0 | 0 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
-120 | 0 | 0.0005 | 0.5 | 0 | 0 | 0 | 0 | 0 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
-110 | 0 | 0.0005 | 0.5 | 0 | 0 | 0 | 0 | 0 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
-100 | 0 | 0.0005 | 0.5 | 0 | 0 | 0 | 0 | 0 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
-90 | 0 | 0.0005 | 0.5 | 0 | 0 | 0 | 0 | 0 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
-80 | 0 | 0.0005 | 0.5 | 0 | 0 | 0 | 0 | 0 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
-70 | 0 | 0.0005 | 0.5 | 0 | 0 | 0 | 0 | 0 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
-60 | 0 | 0.0005 | 0.5 | 0 | 0 | 0 | 0 | 0 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
-50 | 0 | 0.0005 | 0.5 | 0 | 0 | 0 | 0 | 0 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
-40 | 0 | 0.0005 | 0.5 | 0 | 0 | 0 | 0 | 0 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
-30 | 0 | 0.0005 | 0.5 | 0 | 0 | 0 | 0 | 0 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
-20 | 0 | 0.0005 | 0.5 | 0 | 0 | 0 | 0 | 0 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
-10 | 0 | 0.0005 | 0.5 | 0 | 0 | 0 | 0 | 0 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
0 | 0 | 0.0005 | 0.5 | 0 | 0 | 0 | 0 | 0 |  | #DIV/0! | 0 | #DIV/0! | 0 | 0°C
10 | 0 | 0.0005 | 0.5 | 0 | 0 | 0 | 0 | 0 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
20 | 0 | 0.0005 | 0.5 | 0 | 0 | 0 | 0 | 0 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | AMBIENT
30 | 0 | 0.0005 | 0.5 | 0 | 0 | 0 | 0 | 0 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
40 | 0 | 0.0005 | 0.5 | 0 | 0 | 0 | 0 | 0 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
50 | 0 | 0.0005 | 0.5 | 0 | 0 | 0 | 0 | 0 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
60 | 0 | 0.0005 | 0.5 | 0 | 0 | 0 | 0 | 0 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
70 | 0 | 0.0005 | 0.5 | 0 | 0 | 0 | 0 | 0 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
80 | 0 | 0.0005 | 0.5 | 0 | 0 | 0 | 0 | 0 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
90 | 0 | 0.0005 | 0.5 | 0 | 0 | 0 | 0 | 0 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
100 | 0 | 0.0005 | 0.5 | 0 | 0 | 0 | 0 | 0 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
110 | 0 | 0.0005 | 0.5 | 0 | 0 | 0 | 0 | 0 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
120 | 0 | 0.0005 | 0.5 | 0 | 0 | 0 | 0 | 0 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
130 | 0 | 0.0005 | 0.5 | 0 | 0 | 0 | 0 | 0 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
140 | 0 | 0.0005 | 0.5 | 0 | 0 | 0 | 0 | 0 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
150 | 0 | 0.0005 | 0.5 | 0 | 0 | 0 | 0 | 0 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
160 | 0 | 0.0005 | 0.5 | 0 | 0 | 0 | 0 | 0 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
170 | 0 | 0.0005 | 0.5 | 0 | 0 | 0 | 0 | 0 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
180 | 0 | 0.0005 | 0.5 | 0 | 0 | 0 | 0 | 0 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
190 | 0 | 0.0005 | 0.5 | 0 | 0 | 0 | 0 | 0 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
200 | 0 | 0.0005 | 0.5 | 0 | 0 | 0 | 0 | 0 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
210 | 0 | 0.0005 | 0.5 | 0 | 0 | 0 | 0 | 0 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
220 | 0 | 0.0005 | 0.5 | 0 | 0 | 0 | 0 | 0 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
230 | 0 | 0.0005 | 0.5 | 0 | 0 | 0 | 0 | 0 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
240 | 0 | 0.0005 | 0.5 | 0 | 0 | 0 | 0 | 0 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
250 | 0 | 0.0005 | 0.5 | 0 | 0 | 0 | 0 | 0 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
260 | 0 | 0.0005 | 0.5 | 0 | 0 | 0 | 0 | 0 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
270 | 0 | 0.0005 | 0.5 | 0 | 0 | 0 | 0 | 0 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
280 | 0 | 0.0005 | 0.5 | 0 | 0 | 0 | 0 | 0 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
290 | 0 | 0.0005 | 0.5 | 0 | 0 | 0 | 0 | 0 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
300 | 0 | 0.0005 | 0.5 | 0 | 0 | 0 | 0 | 0 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
310 | 0 | 0.0005 | 0.5 | 0 | 0 | 0 | 0 | 0 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
320 | 0 | 0.0005 | 0.5 | 0 | 0 | 0 | 0 | 0 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!

## STD Vals (hidden)
Decade Vals |  | Standard PT Values
4700 | Ω | 100 | Ω
5000 | Ω | 500 | Ω
7500 | Ω | 1000 | Ω
10000 | Ω | Custom | Ω
15000 | Ω
20000 | Ω